# Dual-IMU Knee Goniometer — Design Decisions

Compilation of the design choices, their rationale, and what was done

# About this document

This document collects, in one place, the design decisions made during the dual-IMU knee goniometer proof of concept (August–October 2026): what was decided, why, which alternatives were considered, and where the evidence lives. It is a companion to:

- README.md — the technical reference (math, wire protocol, commands).

- docs/SOP_Build_and_Use.md — how to build and operate the device.

- docs/HANDOVER.md — current state, open issues and next steps.

- docs/areas_to_explore.md — longer analyses of directions not pursued yet.

- analysis/static_test/Static_Stability_Test_Report.docx — the Preliminary Stationary Test Report, with FINDINGS.md as its detailed working record.

Each decision follows the same layout: Decision, Why, Alternatives considered, Evidence (commits / documents), Status. Commit hashes refer to this repository’s history.

# 1. Problem and goals

The goal is a low-cost wearable that measures knee flexion angle from two inertial sensors, one above and one below the knee.

| Goal | What it means in practice |
| --- | --- |
| Placement independence | The boards should not need to be mounted at a precise or repeatable orientation on the limb. |
| Robust signal | Stable at rest, smooth through motion, no drift in the angle. |
| Simple, verifiable PoC | Prefer the smallest system that can be checked on the bench over more capable but unverified machinery. |
| Usable in testing sessions | A GUI that drives calibration, shows live data, and saves every session automatically. |

Scope: sagittal-plane knee flexion (standing ROM, gait, sit-to-stand), relative to a calibration pose. Absolute anatomical angles and 3-D kinematics are out of scope (see §4.5 and §9).

# 2. System architecture

shank board (peripheral)  --UART-->  thigh board (central)  --USB serial-->  PC
   BMI270, Mahony filter   <--keep--    BMI270, Mahony filter                 knee_gui.py / knee_collector_uart.py
                              alive                                           (calibration, angle, CSV)

## 2.1 Two-board central/peripheral topology

- Decision: The thigh board (central) runs its own filter, receives the shank board’s (peripheral) state over a wire, merges both and sends one line per sample to the PC. The PC computes the angle.

- Why: Only one board needs a PC connection; both segments arrive already time-aligned on the central’s clock; the angle math stays in Python where it is easy to test and change without reflashing.

- Alternatives: Each board on its own USB port (needs PC-side time alignment of two streams); computing the angle on the board (harder to test and iterate).

- Evidence: README.md (Hardware & wiring); d9ccd14 renamed master/slave → central/peripheral.

- Status: In use.

## 2.2 Hardware: Arduino Nano 33 BLE Rev2

- Decision: Two Arduino Nano 33 BLE Rev2 boards, using the onboard BMI270 accelerometer + gyroscope. The onboard BMM150 magnetometer is not used.

- Why: Off-the-shelf, inexpensive, IMU on board, Arduino tooling, BLE radio available for a future wireless version.

- Status: In use. One of the two boards has a faulty BMI270 (see §8.1).

## 2.3 Physical build as it stands

- Decision: Each board sits in a basic case. The boards are joined by soldered wires (two UART lines + ground). Each board is powered by its own external USB supply, and the central is connected to the PC over USB serial for data.

- Why: Fastest route to a working, reliable bench prototype.

- Trade-offs: The subject is tethered to the PC and to power; a wire runs across the knee. Battery power and wireless links are future work (§9).

- Status: Bench prototype. No strap or on-body mounting has been finalised.

# 3. Sensor fusion (on each board)

## 3.1 6-DOF Mahony filter, no magnetometer

- Decision: Each board estimates its orientation with a 6-DOF Mahony filter (accelerometer + gyroscope), TWO_KP = 1.0, TWO_KI = 0.04.

- Why: The angle method (§4) only needs the direction of gravity in each board’s frame, which 6-DOF provides without drift. A magnetometer is disturbed by nearby metal (furniture, treadmills, the USB cable, the other board) and needs per-board hard/soft-iron calibration.

- Alternatives: A 9-DOF path with runtime magnetometer calibration and an automatic BMM150→BMI270 axis solver was built (146cc14, 87096d9, d79e51a, 9cf21cb) and then removed (0dd3a93) to get back to a clean, verifiable baseline — it was never confirmed on hardware, not measured and found worse. Analysis in docs/areas_to_explore.md Area 1 concludes it adds nothing for sagittal angles.

- Status: In use. 9-DOF code is recoverable from history if 3-D is ever needed.

## 3.2 Fixing the library’s left-handed axes

- Decision: Negate the x axis of both accelerometer and gyroscope after reading them.

- Why: Arduino_BMI270_BMM150 returns x=-sensor.y, y=-sensor.x, z=sensor.z — a reflection (determinant −1). A quaternion filter needs a right-handed frame; without the fix the gyro’s rotation sense was mirrored and the filter wandered.

- Evidence: 6f1bdc7; README finding #1; comment block at the top of both sketches.

- Status: In use. Any new firmware must keep this.

## 3.3 Gyro bias calibration with a sanity gate

- Decision: At the start of each collection, average ~300 gyro samples (up to 4 s) and subtract the result as bias, but discard it if any axis exceeds BIAS_SANITY_DPS = 3 °/s.

- Why: Bias is the main drift source in a 6-DOF filter. If the board is moving during calibration, the “bias” is a false rate that can make the filter diverge, so it is safer to use zero.

- Evidence: fd0d53b, 6f1bdc7; README finding #2.

- Status: In use. Calibration happens when a collection starts (USB port opened), not at power-up (86efe6d).

## 3.4 Soft accelerometer-trust gate

- Decision: Weight the accelerometer correction by how close |accel| is to 1 g: full trust within 0.10 g (ACC_TRUST_FULL_G), ramping to none at 0.60 g (ACC_TRUST_ZERO_G).

- Why: During brisk motion the accelerometer reads gravity plus linear acceleration. A hard on/off gate let the filter run gyro-only through a fast move, drift, then snap back — the “wrong reading right after returning from a big angle” symptom. A soft ramp always applies some correction.

- Evidence: fd0fbaa (hard gate), 4effe4c (soft gate); README finding #7.

- Status: In use; not yet tuned on dynamic data.

# 4. Knee-angle method (PC side)

## 4.1 Gravity-referenced sagittal inclinometer

- Decision: For each board, take the direction of gravity in the board’s own frame (from the fused quaternion), express it as a signed inclination in that segment’s sagittal plane, and define the knee angle as incl_thigh − incl_shank.

- Why: Gravity-in-board is drift-free in tilt (the accelerometer pins it) and invariant to heading, so the 6-DOF yaw drift never enters. Any constant mounting rotation cancels, which gives placement independence.

- Alternatives: (a) Relative quaternion q_thigh* ⊗ q_shank with swing–twist decomposition — rejected because each board’s yaw drifts independently in 6-DOF, so the angle drifted even on a rigid ruler (README finding #3). (b) Four selectable modes (--method × --gravity-source) — consolidated to one to stop modes masking each other’s failures (4effe4c).

- Evidence: 276d9dc, 4effe4c, 9f5c8c9; README “How it works”; --selftest proves cancellation against random mounts and per-board yaw drift.

- Status: In use. Code: gravity_in_board(), sagittal_inclination(), gravity_knee_angle() in knee_collector_uart.py.

## 4.2 Gravity from the fused quaternion, not the raw accelerometer

- Decision: Read gravity from the Mahony quaternion; keep the raw accelerometer only as a filter-free cross-check.

- Why: The raw accelerometer is corrupted during motion; the fused quaternion lets the gyro carry fast movement while the accelerometer keeps tilt drift-free. (Raw accelerometer was briefly the default while the handedness bug made the quaternion wrong — README finding #4.)

- Status: In use. Raw accelerometer values are still streamed and logged (useful for checking physical movement, as in the stationary report).

## 4.3 Two-step calibration: static zero + functional sweep

- Decision: Each session starts with

  - Zeroing (CAL_SECONDS = 2 s): leg straight and still; the average gravity direction becomes each segment’s 0° axis d_i.

  - Sweep (SWEEP_SECONDS = 6 s): a few slow movements bending knee and hip; the in-plane tilt direction becomes each segment’s forward axis f_i.

- Why: These two vectors define each segment’s sagittal plane in its own board frame without knowing how the board is mounted.

- Detail: Samples tilting less than SWEEP_MIN_ANGLE_DEG = 5° are ignored so noise cannot define the direction. A segment that does not tilt ≥ 5° during the sweep contributes 0 (treated as fixed).

- Consequence observed in testing: In all bench tests only the shank node was swept, so the angle came from the shank alone. Both-node sweeps are a required next step (Stationary Test Report, Limitations).

- Status: In use.

## 4.4 Signed, continuous angle

- Decision: incl = atan2(g·f, g·d): flexion positive, hyperextension negative, continuous through the full range with no ±90° wrap.

- Why: A signed four-quadrant angle distinguishes flexion from hyperextension and retraces identically on the way back from high flexion (verified in --selftest with a 0→130°→0 sweep, no stuck zeros on return).

## 4.5 Relative zero, not an absolute anatomical angle

- Decision: Accept that 0° is “the pose held during zeroing”.

- Why: ROM, rep counting, gait and movement quality are difference-based, so the zero cancels. An absolute angle (e.g. a flexion contracture) needs an external reference that observes the anatomy; the sensors cannot supply it, and a magnetometer does not help (Area 1).

- Alternatives: Functional joint-axis calibration or a known-pose fixture (Area 2) — identified, not prototyped.

- Evidence: 91bc803; README “What 0° means”; docs/areas_to_explore.md Areas 1–2.

## 4.6 Accepted limitation: sagittal plane only

Gravity gives 2 of 3 rotational degrees of freedom; rotation about vertical is unobservable. The method therefore under-reads motion well out of the vertical plane (e.g. lying down with large hip rotation). This is an observability limit, accepted for upright knee flexion.

# 5. Board-to-board link and data protocol

## 5.1 Wired UART at 115 200 baud

- Decision: Serial1 UART at 115 200 baud, both directions wired (peripheral TX → central RX for data; central TX → peripheral RX for the keepalive), common ground.

- Why: Simple and deterministic. The rate was briefly raised to 460 800 to shorten time on the wire, which masked a framing bug and added sensitivity to clock drift as boards warm up; 115 200 gives ~4× the per-byte timing margin and is far above the ~15 kbit/s needed.

- Alternatives: BLE (Area 3) — deferred.

- Evidence: fd0fbaa, c767bc7; README finding #6.

## 5.2 Framed binary packet with resync and checksum

- Decision: 30-byte packet: 0xAA header, quaternion (4 floats), raw accel (3 floats), XOR checksum. The central resyncs on the header and drops packets that fail the checksum.

- Why: The original fixed-length read had no resync, so one lost byte shifted every following packet. Now a glitch costs one packet.

- Evidence: c767bc7.

## 5.3 Peripheral streams; central never waits

- Decision: The peripheral streams packets at 50 Hz; the central drains its UART buffer each loop and keeps the freshest valid packet. If that packet is older than SHANK_STALE_US = 30 ms the shank is reported invalid (zeros) and the PC forward-fills it. The packet age is the line’s last field.

- Why: Polling with a deadline lost samples whenever the peripheral’s RTOS stalled for a few ms; widening the deadline only stalled the central. Streaming decouples the two boards.

- Evidence: 3b97e85, 7197327, 4fc3a07, c111974; README finding #6.

## 5.4 Collect-on-demand

- Decision: Both boards run their IMUs only while the central’s USB port is open. On open, the central calibrates and sends S keepalives every 200 ms; the peripheral idles if none arrives for 1 s, or on X.

- Why: Nothing runs unobserved; each session starts with a fresh bias calibration on a still board, rather than relying on a boot-time one.

- Evidence: 86efe6d.

## 5.5 Fixed 50 Hz output, one write per line

- Decision: Both boards emit at a fixed 50 Hz (EMIT_PERIOD_US = 20 000) on a fixed schedule; the central builds each PC line in memory and sends it with one write. Filters still run at the sensor’s native rate.

- Why: ~104 Hz left little timing margin and 50 Hz is ample for knee ROM. Testing showed the central actually ran at ~43.5 Hz because it printed each line as ~36 blocking prints (~11.7 ms); the one-write line takes ~1–1.8 ms and restored a true 50 Hz.

- Evidence: 8211f4a, 5c8f67d, e02ae6d; Stationary Test Report, Appendix.

## 5.6 Diagnostics channel

- Decision: With DIAG 1 (the current default in both sketches) the peripheral sends a once-a-second health frame (0xAB header) that the central relays as # PDIAG …, and the central adds # CDIAG …. The GUI saves all # lines to knee_diag_<timestamp>.log; analysis/diagnostics/parse_diag.py turns them into tables.

- Why: Needed to find the sample-rate shortfall and the faulty sensor.

- Constraint: The PDiag struct must match byte for byte in both sketches; flash both boards from the same commit.

- Evidence: 72943ba.

# 6. Robustness and fault handling

## 6.1 Peripheral self-healing

- Low-rate start: if fewer than 30 samples arrive in 1 s, re-initialise the sensor (the BMI270 occasionally starts at a wrong ~10 Hz rate) — 4036498.

- Data-ready stall: if no sample for 250 ms, LED solid and re-init every 500 ms — 5f8aaff.

- IMU init failure at boot: fast-blink ~3 s, then reboot.

## 6.2 Sensor-fault guard

- Decision: On the peripheral, reject a sample if the read took > 50 ms, |accel| is < 0.25 g or > 4 g, or the gyro reads > 1000 °/s while |accel| is within 0.15 g of 1 g. Rejected samples are sent as an all-zero quaternion (a gap, not a wrong angle). Each re-init resets the filter’s integral term. If no good sample arrives for 5 s, the board reboots.

- Why: In endurance runs one board’s BMI270 began returning garbage after 20–38 min and then stopped entirely; this turns wrong angles into visible gaps.

- Note: Only the peripheral has this guard. The central does not.

- Evidence: e02ae6d; Stationary Test Report, “Cause: one faulty board”.

## 6.3 LED indicators

| Board | Pattern | Meaning |
| --- | --- | --- |
| Central | On | Collecting (USB port open) |
| Central | Off | Idle |
| Peripheral | 3 quick flashes | Boot |
| Peripheral | Slow blink | Idle, waiting for keepalive |
| Peripheral | Fast blink (~3 Hz) | Active, streaming |
| Peripheral | Solid on | Sensor stalled (re-init firing) |
| Peripheral | Off / frozen | Unpowered or hard fault |

# 7. PC software

## 7.1 CLI collector plus a live GUI

- Decision: knee_collector_uart.py holds the math, calibration, CSV logic and --selftest; knee_gui.py (Tkinter + Matplotlib) reuses it and adds port scanning, a colour-coded calibration banner, live plots, and auto-saved sessions.

- Why: One tested implementation of the math; a GUI makes test sessions repeatable for non-developers.

- Evidence: 9ef5550, d15c4db, 1a79b4d, a0e852a.

## 7.2 Session-based collect / stop

- Decision: Each “Calibrate & collect” starts a new session with fresh calibration and a new timestamped CSV; “Stop collecting” freezes the plot on the full session for inspection. No pause/resume.

- Why: Avoids mixing calibrations in one file; makes every file self-contained.

## 7.3 One CSV row per received sample

- Decision: Write exactly one row per line from the central, timed by the board’s clock (t_session_s), plus the PC receive time (t_wall_iso).

- Why: The GUI used to write rows on its own 20 ms timer, which ticks unevenly on Windows: ~20 % of rows repeated a sample and others were skipped.

- Evidence: 8a9aaf4; Stationary Test Report, Appendix.

## 7.4 Dropouts: forward-fill vs gap

- Decision: Short gaps (≤ MAX_FILL = 10 samples) are forward-filled; longer gaps are marked missing. The CSV records the status of every sample (valid / filled / missing), and the GUI can display either a filled line or discrete points.

- Why: Keeps plots readable without hiding where data was real.

## 7.5 Clear error diagnosis

diagnose_stream() names the failing link: no data on the port, wrong/old firmware (not an 18-field D line), or a dead peripheral (all shank samples invalid). The CLI and GUI share it.

# 8. Testing and validation

## 8.1 Preliminary stationary testing (Sept 2026)

Full results: analysis/static_test/Static_Stability_Test_Report.docx.

| Test | Result |
| --- | --- |
| Stability (6 × ~5 min holds) | ~0.01–0.02° noise; no noise growth (Allan deviation flat 0.1–60 s) |
| Endurance (24–60 min) | Stable for 33 min on healthy hardware; ≈ −0.001 °/min drift |
| 45° / 90° rigs (preliminary) | Read 45.0° and 89.0° |
| Sample rate | ~40 Hz found; fixed to a true 50 Hz |
| Data loss in long runs | Traced to a faulty BMI270 on one board; the other ran > 1 h clean in either role |

Key decisions from testing:

- Exclude the first 20 s of each session from stationary analysis (calibration and settling).

- Board-swap test to separate a hardware fault from a firmware fault — the fault moved with the board.

- Raw accelerometer as an independent check of physical movement, which showed that Session 2’s creep was the rig settling, not sensor drift.

## 8.2 Earlier bench checks

Rigid-ruler test (both boards on one ruler, true relative angle constant) was the key diagnostic for drift; the right-angle placement read ~90°. Math is covered by knee_collector_uart.py --selftest and knee_gui.py --selftest.

## 8.3 Not yet tested

Both-node calibration; dynamic movement; on-body use; controlled accuracy against a goniometer or precise rig; noise below the 0.01° log resolution.

# 9. Directions considered but not pursued

Detailed in docs/areas_to_explore.md.

| Area | Recommendation |
| --- | --- |
| 1. Magnetometer fusion | Do not add for sagittal angles; only for a true 3-D requirement, after checking the magnetic environment. |
| 2. Absolute anatomical angle | Functional joint-axis calibration or a known-pose reference; not prototyped. |
| 3. BLE (wireless) links | Start with the board-to-PC link; needs source timestamps and batteries. Main benefit is safety / REB approval. |
| 4. More nodes / joints | Needs node IDs, a generalised data model and cross-node time sync; BLE suits a star topology. |

# 10. Timeline of major milestones

| Date (2026) | Milestone | Commits |
| --- | --- | --- |
| Aug 14 | Initial UART master/slave prototype (pitch-based complementary filter) | 09f3557, 4b2fd42 |
| Aug 17 | Quaternion, placement-independent angle; 9-DOF magnetometer path tried | 557ba29, 146cc14 |
| Aug 18 | Strip back to clean 6-DOF; handedness fix; gyro bias; gravity-referenced angle | 0dd3a93, 6f1bdc7, 276d9dc |
| Aug 19 | One angle method; live GUI | 4effe4c, 9ef5550 |
| Aug 20 | Link fixes (resync, stream not poll); collect-on-demand; 50 Hz | c767bc7, 7197327, 86efe6d |
| Sep 21–22 | Central/peripheral rename; research directions documented | d9ccd14, 170580c |
| Sep 23–24 | Stationary testing; diagnostics; rate fix; sensor-fault guard | 99a2446, 72943ba, e02ae6d, 8a9aaf4 |
| Sep–Oct | Preliminary Stationary Test Report | 70815cf, 38724bf |