# Dual-IMU Knee Goniometer — Handover Summary

# Handover summary

State of the project at transition (October 2026), what is in the repository, what is missing, and what to do next.

## Current state

- Working bench prototype. Two Arduino Nano 33 BLE Rev2 boards (thigh = central, shank = peripheral) in basic cases, joined by soldered UART wires, each powered over USB; the central streams to a PC at a true 50 Hz.

- Software complete for PoC testing. Live GUI with guided calibration and auto-saved sessions; CLI collector; self-tests pass.

- Preliminary stationary testing done. ≈0.01–0.02° noise, ≈ −0.001 °/min drift on healthy hardware; 45°/90° rigs read 45.0°/89.0° (preliminary).

- Not yet done: dynamic / on-body testing, both-node calibration, a controlled accuracy test.

## Documents

| Document | Use it for |
| --- | --- |
| docs/SOP_Build_and_Use.md (+ .docx) | Building the device and running a session |
| docs/DESIGN_DECISIONS.md (+ .docx) | Why the system is the way it is |
| analysis/static_test/Static_Stability_Test_Report.docx | Preliminary Stationary Test Report |
| analysis/static_test/FINDINGS.md | Detailed working record behind the report |
| README.md | Technical reference: math, wire protocol, commands |
| docs/areas_to_explore.md | Analyses of magnetometer, absolute angle, BLE, multi-node |

## Repository map

central_imu/central_imu.ino          thigh board firmware
peripheral_imu/peripheral_imu.ino    shank board firmware
knee_collector_uart.py               angle math, calibration, CLI collector, --selftest
knee_gui.py                          live GUI (uses knee_collector_uart.py)
requirements.txt                     Python dependencies
analysis/static_test/                stationary test scripts, figures, report
analysis/diagnostics/parse_diag.py   diagnostic log -> tables
docs/                                SOP, design decisions, handover, research areas
docs/build_docs.sh                   regenerates the .docx files from Markdown

## Open issues

- Faulty board. One board’s BMI270 fails after 20–40 min. It must be identified, labelled and replaced. [TO FILL: which board]

- Raw test data not preserved. The stationary-test CSVs and diagnostic logs were not archived (*.csv / *.log are git-ignored), so the figures in the report cannot be regenerated. Set up a data store before the next test.

- Physical build details not recorded. Case model, wire gauge/length, strain relief and a leg-mounting method are placeholders ([TO FILL]) in the SOP.

- Analysis scripts are working scripts. analysis/static_test/stats.py, plot.py and calibration.py were written ad hoc for this dataset (they read every CSV in the current folder). plot_report.py and plot_angle_rig.py produced the report figures and expect specific file names.

- Peripheral-only fault guard. The sensor-fault guard runs on the peripheral only; the central has no equivalent.

- Angle log resolution. Angles are written to 0.01°, which limits how low a noise floor can be measured.

## Recommended next steps (in order)

- Replace the faulty board; repeat a one-hour stationary run on two healthy boards for a clean baseline.

- Sweep both nodes during calibration so the angle uses both sensors; fix the boards and cables firmly.

- Define and document an on-body mounting method; complete the SOP placeholders.

- Controlled accuracy test: repeated 0°/45°/90° placements on a precise rig or against a goniometer.

- Dynamic range-of-motion testing (tune TWO_KP and the soft accel gate if fast reps lag or overshoot).

- Consider the longer-term directions in docs/areas_to_explore.md (battery + BLE for an untethered device; absolute angle via functional calibration).

## Housekeeping

- No licence file — decide whether one is needed before sharing the code.

- Tag the handed-over commit (e.g. v0.1-handover) so the firmware, software and documents can be referenced together.

- Regenerate the Word copies after editing the Markdown: bash docs/build_docs.sh.